## Slide 1: Detector development and characterisation for neutron- and gamma-based radiography and spectroscopy for NDT
J. Peñas, C. Cabo-Landeira, B. Fernández, C. Guerrero, M.A. Millán-Callado, Y. Ayyad
3rd DONES Users Workshop
Zagreb – October 2024

## Slide 2: Outline
Neutrons for Non-Destructive Testing (NDT).
Neutron and gamma imaging. Detectors description and validation.
Gaseous detector for neutron spectroscopy. Overview, features and performance.
Prospects at IFMIF-DONES facility.
Conclusions.
J.Peñas | 3rd DONES Users Workshop | Zagreb | October 2024

## Slide 3: Non-Destructive Testing
Conventional NDT techniques involve X-rays and gamma rays for transmission radiography. Several limitations:
Dense and/or high-atomic number materials.
Thick objects.
NDT with neutrons constitute a valuable technique:
Low interaction with heavy nuclei.
Enhanced absorption of light elements and others (B, Gd, Cd,...).
Not limited to radiography: PGAA, BETI, FNAA,...
Other properties of matter: composition, internal stress or inclusions, thickness, etc, of interest in NDT.
Several applications in industry: MAM, quality control, pipeline maintenance, etc.
X-rays
Gamma
(137Cs)
Thermal neutrons
J.Peñas | 3rd DONES Users Workshop | Zagreb | October 2024

## Slide 4: Gamma and neutron imaging - Detector
Mirror
Scintillator
Lens
Pb & HDPE shielding
Sample
Black Box
CMOS
camera
Radiation
Most common design based on a scintillator and a camera.
  Frame-based cameras (CCD/CMOS).
  Event-driven camera (Timepix).
Flexible to different types of radiation (neutrons or photons) by exchanging the scintillator.
  CsI or NaI for gamma rays.
  ZnS:Cu for neutrons. Polypropylene layer for n-p conversion.
  GADOX for both radiation types.
Optional image intensifier for low flux conditions.​
Rotary stage for tomographic measurements.
J.Peñas | 3rd DONES Users Workshop | Zagreb | October 2024

## Slide 5: Gamma and neutron imaging
Gamma radiography
60Co gamma irradiator (~47 TBq) at USC, Santiago de Compostela.
  AECL Theatron 780 therapy unit.
  Adjustable source size (lead plates) and beam directionality (rotary axis).
Device specifications:
  GADOX scintillating sheet, 10x20 cm area.
  Hamamatsu Orca-Flash4.0 v2 CMOS camera: 2048x2048 pixels with 6.5x6.5 µm pixel size. Optical lens f/1.3.
J.Peñas | 3rd DONES Users Workshop | Zagreb | October 2024
Neutron radiography
HiSPANoS neutron beamline at Centro Nacional de Aceleradores (CNA), Seville.
  Proton or deuterium beams delivered by 3MV tandem accelerator.
  Fast neutrons through (p,n) and (d,n) reactions on a Be target. Flux: ~5·106 n/s/cm2 above 1.3 MeV @1.5 m.
  Moderation with HDPE blocks down to thermal energies.
Device specifications:
  ZnS:Cu scintillator, 10x12 cm area and 2 mm thickness.
  ASIO ZWO ASI533MM CMOS camera: 3000x3000 pixels with 3.7x3.7 um pixel size. Optical lens f/1.8.

## Slide 6
Gamma radiography
Source size: 7x7 cm
Exposure time: 5x7 s
Flux: 5·108 ph/s/cm2
MAM steel sample.
2.5 cm thickness.
Holes with 0.5 - 1.5 mm diameter.​
Source-object dist.: ~80 cm​
Spatial resolution (edges): 0.9 mm
20 – 36 projections.
1.5 mm
J.Peñas | 3rd DONES Users Workshop | Zagreb | October 2024

## Slide 7
Gamma radiography
Source size: 7x7 cm
Exposure time: 5x7 s
Flux: 5·108 ph/s/cm2
MAM steel sample.
2.5 cm thickness.
Holes with 0.5 - 1.5 mm diameter.​
Source-object dist.: ~80 cm​
Spatial resolution (edges): 0.9 mm
20 – 36 projections.
J.Peñas | 3rd DONES Users Workshop | Zagreb | October 2024

## Slide 8
Neutron radiography (thermal)
Highly neutron- absorbing material
Exposure time: 100x10 s
Source-object dist.: ~80 cm
Spatial resolution (edges): 1.4 mm
1.0 mm
Gamma rays
J.Peñas | 3rd DONES Users Workshop | Zagreb | October 2024

## Slide 9
Neutron radiography (thermal)
Lead thickness: 2.5 cm
Lead thickness: 5 cm
1.0 mm
J.Peñas | 3rd DONES Users Workshop | Zagreb | October 2024
Exposure time: 100x10 s
Source-object dist.: ~80 cm
Spatial resolution (edges): 1.4 mm
Highly valuable for NDT inspection!

## Slide 10: Event-driven image adquisition
Timepix3 detector
Array of 256x256 pixels of 5.5x5.5 µm size. Each pixel constitutes a separate acquisition with its own electronics.
Provides additional timing information associated to each singular event:
  (x,y)-coordinates.
  Time-of-arrival (ToA).
  Time-over-threshold (ToT).
Intrinsic temporal resolution of 1.56 ns.
Silicon sensor sensitive to photons and charged particles. Adaptable to neutrons through a conversion layer (polyethylene, paraffin, etc) and/or detector (MCP, GEMs,...).
J.Peñas | 3rd DONES Users Workshop | Zagreb | October 2024

## Slide 11: Event-driven image adquisition
Main advantages over frame-based imaging:
Better image reconstruction methods: improved spatial resolution and noise reduction.
Pulsed sources: relevant timing information derived from ToF techniques.
For neutrons, energy-resolved neutron imaging. Fast, thermal, or at resonance band.
Pulse Shape Discrimination (PSD) on each signal, allowing for simultaneous imaging using different radiation types.
J.Peñas | 3rd DONES Users Workshop | Zagreb | October 2024

## Slide 12: Prospects at the IFMIF-DONES facility
J.Peñas | 3rd DONES Users Workshop | Zagreb | October 2024
Complementary Experiments Hall
R160
(J.Hirtz et al., 2022)
40 MeV deuterons to <55 MeV neutrons. Peak at about 15 MeV.
Fast neutron flux: 1010 n/s/cm2.
Thermal neutron flux:
  107 n/s/cm2 (no collim.)
  104 n/s/cm2 (L/D = 200)
(20 – 25% fast-to-thermal ratio after moderation)
TOF-DONES
R026
(D.Cano-Ott talk @ IFMIF-DONES Workshop, 2022)
40 MeV deuterons to <55 MeV neutrons. Peak at about 15 MeV.
Deuteron pulses of 5.6 ns evenly spaced 5.7 us (1:1000) ratio -> Up to 175 kHz.
Total neutron flux: 106 n/s/cm2 at 22 m ToF distance.

## Slide 13: Prospects at the IFMIF-DONES facility
J.Peñas | 3rd DONES Users Workshop | Zagreb | October 2024
Neutron imaging
High fast neutron flux
High fast neutron energy
High spatial resolution and Signal-to-Noise Ratio (SNR). Shorter exposure times.​
Deep penetration on thick and dense objects.
1010 n/s/cm2@R160, orders of magnitude higher than @CNA
Energy-selective imaging (Timepix, LEELA) at TOF-DONES provides additional information:
Improved conventional imaging due to beam properties:
Bragg-edge transmission imaging (BETI)
Fast Neutron Resonance Radiography (FNRR)
Composition and stress map of the sample (>103 n/s/cm2).
Composition of the sample.
Material-specific resonances at fast neutron energies
Attenuation of low-energy neutrons due to cryst. lattice.
PSD processing (Timepix) allows for radiation type specific imaging.

## Slide 14: Prospects at the IFMIF-DONES facility
J.Peñas | 3rd DONES Users Workshop | Zagreb | October 2024
Neutron flux:
PGAA ~108 n/s/cm2
NAA ~1012-14 n/s/cm2
FNAA ~108 n/s/cm2
Complementary techniques – Material characterisation
Neutron spectroscopy (Timepix, LEELA) at TOF-DONES enable several techniques:
Neutron Resonance Transmission Analysis (NRTA)
Neutron back-diffusion
Decay gammas after activation
Prompt Gamma Activations Analysis (PGAA)​
Gammas from fast neutron inelastic scattering​
Neutron Activation Analysis (NAA)
Fast Neutron Activation Analysis (FNAA)
Other techniques based on gamma spectroscopy also available (Timepix):
Detector calibration and characterisation
LEELA: efficiency, light yield, energy resolution, etc.
Deexcitation gammas
Thickness of the sample.

## Slide 15
Neutron sources are valuable for NDT.
Frame-based and event-driven (Timepix) cameras.
Gaseous spectroscopic detector based on stacks of O-PPACs.
Several techniques and experiments available at IFMIF-DONES facility!
Conclusions
J.Peñas | 3rd DONES Users Workshop | Zagreb | October 2024

## Slide 16: Thank you for your attention!

## Slide 17: Back-up

## Slide 18: Summary
J.Peñas | 3rd DONES Users Workshop | Zagreb | October 2024
Gaseous spectroscopic detector
Frame-based imaging system
Event-driven imaging system
Complementary Experiments Hall R160
TOF-DONES R026
High resolution thermal and fast imaging
All of the above
High resolution thermal and fast imaging
Energy-resolved neutron imaging
  BETI
  FNRR
All of the above
ToF imaging
NAA, FNAA, PGAA
NRTA
Neutron back-diffussion
High resolution thermal and fast imaging
Energy-resolved neutron imaging
  BETI
  FNRR
All of the above
NRTA
Neutron back-diffussion
Imaging
Material char.
Both

## Slide 19: Other results from CNA
Thermal neutrons
Fast neutrons
J.Peñas | 3rd DONES Users Workshop | Zagreb | October 2024